## Slide 1: Optical Coherence Tomography (OCT): Image-guided Surgery of Breast Cancer
Yiwen Meng

## Slide 2: Breast Cancer has significant impact on human health
211,300 new cases of invasive breast cancer and 55700 new cases of in situ breast cancer in 2003
The most frequently diagnosed non-skin cancer in women, ranks second among cancer deaths in women (second to lung cancer) with an estimated 40,200 deaths anticipated in 2003
Early stage breast cancer and other breast malignancies are often difficult to detect and track using mammograms

## Slide 3: Current Detection Methods and Treatment Protocols Must still be Improved
5-year relative survival rate for localized breast cancer is now at 97%, compared to only 72% in the 1940s
5-year survival rate falls to 78% because of loco-regional metastasis and only 23 % for distant metastasis
After 5 years, survival following the diagnosis of breast cancer continues to fall for all stages

## Slide 4: Surgical Treatment of Breast Cancer still needs to be improved
Approximately 40–50% of local tumors develop recurrence following optimum, surgical or radiation therapy.
The major reason: No relatively accurate definition of tumor margin
The lack of well-developed technology: distinguish tumor margin
MRI, Ultrasound: mm resolution, not μm (microscopic spread of malignant cells)

## Slide 5: Surgical Treatment of Breast Cancer still needs to be improved
Current Oncology: excise tumor with an “arbitrary margin” and send the specimen to assess the margins
Radiation therapy: the entire breast even if the margins are presumed to be free of occult disease
Common surgical practice: take wide surgical margins of primary tumor
Detrimental for tumors located in the brain sensitive and functional regions (Brain) => minimize surgical removal of normal tissue
Image-guided interventional system capable of scanning both the resected tissue and the walls of the tumor cavity for neoplastic tissue or cells in real-time and at cellular resolutions

## Slide 6: Radio-Guided Surgery
Radio-label a cancer-specific monoclonal antibody (MAb) to bind tumor
A handheld counter is used to define tumor margins (which show radioactivity)
Not possible to prepare MAbs for all solid tumors
Most MAbs are hybrids of mouse proteins => foreign protein allergy
The accuracy of detecting cancerous tissues is not as efficient as was originally presumed

## Slide 7: Diffuse optical tomography (DOT) for Breast Caner Detection
Near-infrared light is capable of propagating deeply through biological tissue
An alternative to X-ray that uses ionizing radiation
Successful images of tumors and contrast-agent tagged sites for large tumor (~1 cm) in breast
Powerful for invasive scenario: fine-needle aspirations, core-needle biopsies, or surgical resections
Lost spatial information and imaging resolution: multiple random scattering events through highly-scattering tissue

## Slide 8: Principle of DOT

## Slide 9: Characteristics of OCT System
Near-infrared wavelengths, analogous to ultrasound
Resolution: few mm for highly- scattering tissue, less than 1 μm
Generates cross-sectional images representing the spatially-localized intensity of optical backscatter of light passed through tissue
Powerful alternative to standard histological processing in the investigation and treatment of breast cancer
Rapidly scan large areas of tissue for suspicious breast architecture or morphology to guide at the cellular level and the surgical resection of neoplastic disease, and to scan tumor margins

## Slide 10: Characteristics of OCT System
OCT imaging can be performed through small needles and hand-held probes
Image both in vitro and in vivo tissues to identify regions that suggest abnormalities and should be biopsied for histopathological examination
Utilizes a technique known as low-coherence interferometry
Optical interference occurs when the optical distance traveled by the light in the tissue sample and reference paths match to within the coherence length of the light source

## Slide 11: Schematics of OCT System

## Slide 12: Configuration of OCT System
Laser output wavelength: center 800nm with a 20 nm bandwidth
Ultrahigh numerical aperture fiber (UHNA): broaden the light and improve axial imaging resolution to 2 μm
Two-dimensional images (512*512 pixels) were acquired in 17 s
Three-dimensional data sets consisting of 60 2-D images required 17min
Faster acquisition rates are feasible (125μs per image, 7.5 s per 3-D data set) at the expense of lower signal-to-noise ratios

## Slide 13: Beam-Delivery System Of OCT
Surgical or research microscope
Fine-needle aspiration,
core-needle biopsy probes

## Slide 14: Tissue and Image Analysis between OCT and Light Microscope
Light microscopy :Olympus BH-2 microscope for digital images
Comparisons were made on images correlating microstructural features
Three-dimensional data sets were viewed and analyzed using public-domain freeware ImageJ
The OCT and histological images presented represent the best match based on the observed architectural morphology

## Slide 15: OCT Images of late-stage Tumor
OCT
Light Microscopy
Normal
Tumor
Fat (Adipocytes) :
Low Scattering
Tumor:
High Scattering

## Slide 16: OCT Images of Early-stage Tumor
OCT
Light Microscopy
Normal
Tumor
Box: regions of interest for corresponding histology

## Slide 17: OCT of Cellular Features along Tumor Margin
(A): OCT image of a tumor margin
(B) and (C): corresponding histopathology
Arrow and box: migrating metastatic cells intra-operatively

## Slide 18: Three-Dimensional OCT Images
Normal
Tumor
A series of 60 2-D OCT images were acquired and assembled for each projection, 50 μm inter-slice spacing.

## Slide 19: Potential of OCT for Identifying Breast Cancer in vivo
A large portion of breast is composed of fat (adipocytes): low scattering; Tumor cells: high-scattering.
Resolution(2 μm axial, 10 μm transverse) enable the visualization of larger cells such as adipocytes and undifferentiated cells (20–40 μm)
Molecular imaging techniques enables OCT by the ability to design and target contrast agents with molecular specificity
Enabling OCT to Clearly differentiate normal from abnormal breast tissue at the cellular or molecular level

## Slide 20: OCT for Research and Intra-operative Surgical Diagnostics
OCT image-guided placement and assessment of retinal laser lesions in the in vivo primate eye
Examined the dynamic evolution of retinal lesions from continuous wave lasers at visible wavelengths
Images were acquired every 5 s, the time evolution of lesion morphology enabled a better understanding of the mechanisms of damage
Faster image acquisition rates: enable rapid microstructural changes to be observed sequentially over time
Guide and assess the progress of surgical interventions using high-power argon laser ablation and radio-frequency (RF) ablation

## Slide 21: OCT Makes a significant Impact on Research and Treatment
High-resolution: mm to 1 μm
Deep penetration through tissue: longer wavelengths is scattered less (1300 nm), increase imaging depths by 2 to 3 times, lower resolution
Three-dimension image reconstruction
Enable these histopathological observations and diagnoses to be made intraoperatively in real-time
OCT image-guided fine-needle aspirations and core-needle biopsies

## Slide 22: Thank you! Questions?

## Slide 23: Appendix I: OCT System
The sample arm light from the fiber was collimated, reflected off a pair of orthogonal galvanometer-mounted mirrors that directed the X –Y position of the beam on the tissue, and was focused by a 20mm focal length achromatic lens to a spot size diameter (transverse resolution) of 10 μm
The reference arm contained a galvanometer-mounted retroreflecting mirror translated over an optical path length of 3mm at a frequency of 30 Hz
Reflections from each arm were combined in the fiber-optic coupler and sent to a photodiode.
The electrical output of the photodiode was bandpass filtered, amplified, and digitized with 12-bit accuracy.

## Slide 24: Appendix II: Breast Cancer Research Model
Female Sprague–Dawley rats
Induction of mammary tumors by injection with the direct-acting carcinogen N-methyl-N-nitrosourea (MNU)
MNU is the most widely-used animal model for studying breast cancer development
Mammary carcinogenesis in this model mimic numerous characteristics associated with human breast carcinogenesis